## spe_30_header_format3
Binary | Type_Name | Offset | Description | Versions tested | struct format | Content
16s | ControllerVersion | 0 | Hardware Versionx
16s | LogicOutput | 2 | Definition of Output BNC
16u | AmpHiCapLowNoise | 4 | Amp Switching Mode
16u | xDimDet | 6 | Detector x dimension of chip.
16s | mode | 8 | timing mode
32f | exp_sec | 10 | alternative exposure | 2.2 | f | Exposuretime in sec
16s | VChipXdim | 14 | Vitual Chip X dim
16s | VChipYdim | 16 | Virtual Chip Y dim
16u | yDimDet | 18 | y dimension of CCD or detector.
8s | date[DATEMAX] | 20 | datex | 2.2 | 16s | Date
16s | VirtualChipFlag | 30 | On/Off
8s | Spare_1[2] | 32
16s | noscan | 34 | Old number of scans - should always be -1
32f | DetTemperature | 36 | Detector Temperature Setx
16s | DetType | 40 | CCD/DiodeArray type
16u | xdim | 42 | actual of pixels on x axisx | 2.2, 3.0 | H | Frame Width
16s | stdiode | 44 | trigger diode
32f | DelayTime | 46 | Used with Async Mode
16u | ShutterControl | 50 | Normal
16s | AbsorbLive | 52 | On/Off
16u | AbsorbMode | 54 | Reference Strip or File
16s | CanDoVirtualChipFlag | 56 | T/F Cont/Chip able to do Virtual Chip
16s | ThresholdMinLive | 58 | On/Off
32f | ThresholdMinVal | 60 | Threshold Minimum Value
16s | ThresholdMaxLive | 64 | On/Off
32f | ThresholdMaxVal | 66 | Threshold Maximum Value
16s | SpecAutoSpectroMode | 70 | T/F Spectrograph Used
32f | SpecCenterWlNm | 72 | Center Wavelength in Nm | 2.2 | f | Central Wavelength in nm
16s | SpecGlueFlag | 76 | T/F File Is Glued
32f | SpecGlueStartWlNm | 78 | Starting Wavelength in Nm
32f | SpecGlueEndWlNm | 82 | Ending Wavelength in Nm
32f | SpecGlueMinOvrlpNm | 86 | Minimum Overlap in Nm
32f | SpecGlueFinalResNm | 90 | Final Resolution in Nm
16s | PulserType | 94 | 0=None
16s | CustomChipFlag | 96 | T/F Custom Chip Used
16s | XPrePixels | 98 | Pre Pixels in X direction
16s | XPostPixels | 100 | Post Pixels in X direction
16s | YPrePixels | 102 | Pre Pixels in Y direction
16s | YPostPixels | 104 | Post Pixels in Y direction
16s | asynen | 106 | asynchron enable flag 0 = off
16s | datatype | 108 | experiment datatype: 0 = 32f (4 bytes) | 2.2, 3.0 | h | Datatype
16s | PulserMode | 110 | Repetitive/Sequential
16u | PulserOnChipAccums | 112 | Num PTG On-Chip Accums
32u | PulserRepeatExp | 114 | Num Exp Repeats (Pulser SW Accum)
32f | PulseRepWidth | 118 | Num Exp Repeats (Pulser SW Accum) (usec)
32f | PulseRepDelay | 122 | Width Value for Repetitive pulse (usec)
32f | PulseSeqStartWidth | 126 | Start Width for Sequential pulse (usec)
32f | PulseSeqEndWidth | 130 | End Width for Sequential pulse (usec)
32f | PulseSeqStartDelay | 134 | Start Delay for Sequential pulse (usec)
32f | PulseSeqEndDelay | 138 | End Delay for Sequential pulse (usec)
16s | PulseSeqIncMode | 142 | Increments: 1=Fixed
16s | PImaxUsed | 144 | PI-Max type controller flag
16s | PImaxMode | 146 | PI-Max mode
16s | PImaxGain | 148 | PI-Max Gain
16s | BackGrndApplied | 150 | 1 if background subtraction donex | 2.2 | i | 1 if Background Substraction is done
16s | PImax2nsBrdUsed | 152 | T/F PI-Max 2ns Board Used
16u | minblk | 154 | min. # of strips per skips
16u | numminblk | 156 | # of min-blocks before geo skps
16s | SpecMirrorLocation[2] | 158 | Spectro Mirror Location
16s | SpecSlitLocation[4] | 162 | Spectro Slit Location
16s | CustomTimingFlag | 170 | T/F Custom Timing Used
8s | ExperimentTimeLocal[TIMEMAX] | 172 | Experiment Local Time as hhmmss | 2.2 | 6s | Local Timestamp hhmmss
8s | ExperimentTimeUTC[TIMEMAX] | 179 | Experiment UTC Time as hhmmss | 2.2 | 16s | UTC Timestamp hhmmss
16s | ExposUnits | 186 | User Units for Exposure
16u | ADCoffset | 188 | ADC offset
16u | ADCrate | 190 | ADC rate
16u | ADCtype | 192 | ADC type
16u | ADCresolution | 194 | ADC resolution
16u | ADCbitAdjust | 196 | ADC bit adjust
16u | gain | 198 | gain
8s | Comments[5][COMMENTMAX] | 200 | File Comments
16u | geometric | 600 | geometric ops: rotate 0x01
8s | xlabel[LABELMAX] | 602 | intensity display stringx
16u | cleans | 618 | cleans
16u | NumSkpPerCln | 620 | number of skips per clean.
16s | SpecMirrorPos[2] | 622 | Spectrograph Mirror Positionsx
32f | SpecSlitPos[4] | 626 | Spectrograph Slit Positions
16s | AutoCleansActive | 642 | T/F
16s | UseContCleansInst | 644 | T/F
16s | AbsorbStripNum | 646 | Absorbance Strip Numberx
16s | SpecSlitPosUnits | 648 | Spectrograph Slit Position Unitsx
32f | SpecGrooves | 650 | Spectrograph Grating Grooves | 2.2 | 32f | Grating
16s | srccmp | 654 | number of source comp.diodes
16u | ydim | 656 | y dimension of raw data | 2.2, 3.0 | H | Frame Height
16s | scramble | 658 | 0=scrambled
16s | ContinuousCleansFlag | 660 | T/F Continuous Cleans Timing Option
16s | ExternalTriggerFlag | 662 | T/F External Trigger Timing Option
32s | lnoscan | 664 | Number of scans (Early WinX)
32s | lavgexp | 668 | Number of Accumulationsx
32f | ReadoutTime | 672 | Experiment readout time
16s | TriggeredModeFlag | 676 | T/F Triggered Timing Option
64u | XMLOffset | 678 | Starting location of the XML footer | 3.0 | 64Q | Offset of the XML Footer
8s | sw_version[FILEVERMAX] | 688 | Version of SW creating this file
16s | type | 704 | 1 = new120 (Type II)
16s | flatFieldApplied | 706 | 1 if flat field was applied
8s | Spare_3[16] | 708
16s | kin_trig_mode | 724 | Kinetics Trigger Mode
8s | dlabel[LABELMAX] | 726 | Data label
8s | Spare_4[436] | 742
8s | PulseFileName[HDRNAMEMAX] | 1178 | Name of Pulser File with Pulse Widths/Delays (for Z-Slice)
8s | AbsorbFileName[HDRNAMEMAX] | 1298 | Name of Absorbance File (if File Mode)
32u | NumExpRepeats | 1418 | Number of Times experiment repeated
32u | NumExpAccums | 1422 | Number of Time experiment accumulated
16s | YT_Flag | 1426 | Set to 1 if this file contains YT data
32f | clkspd_us | 1428 | Vert Clock Speed in micro-sec
16s | HWaccumFlag | 1432 | set to 1 if accum done by Hardware
16s | StoreSync | 1434 | set to 1 if store sync used
16s | BlemishApplied | 1436 | set to 1 if blemish removal applied
16s | CosmicApplied | 1438 | set to 1 if cosmic ray removal applied
16s | CosmicType | 1440 | if cosmic ray applied
32f | CosmicThreshold | 1442 | Threshold of cosmic ray removal
32s | NumFrames | 1446 | number of frames in file | 2.2, 3.0 | i | Number of Frames in File
32f | MaxIntensity | 1450 | max intensity of data (future)
32f | MinIntensity | 1454 | min intensity of data future)
8s | ylabel[LABELMAX] | 1458 | y axis label
16u | ShutterType | 1474 | shutter type
32f | shutterComp | 1476 | shutter compensation time
16u | readoutMode | 1480 | readout mode
16u | WindowSize | 1482 | window size for kinetics only
16u | clkspd | 1484 | clock speed for kinetics & frame transfer
16u | interface_type | 1486 | computer interface (isa-taxi
16s | NumROIsInExperiment | 1488 | May be more than the 10 allowed in this header (if 0)
8s | Spare_5[16] | 1490
16u | controllerNum | 1506 | if multiple controller system will have controller number data came from. This is a future item.
16u | SWmade | 1508 | Which software package created this file
16s | NumROI | 1510 | number of ROIs used. if 0 assume 1
16u | ROI1_startx | 1512 | ROI1 left x start value
16u | ROI1_endx | 1514 | ROI1 right x value
16u | ROI1_groupx | 1516 | ROI1 amount x is binned/grouped in hw
16u | ROI1_starty | 1518 | ROI1 top y start value
16u | ROI1_endy | 1520 | ROI1 bottom y value
16u | ROI1_groupy | 1522 | ROI1 amount y is binned/grouped in hw
16u | ROI2_startx | 1524 | ROI2 left x start value
16u | ROI2_endx | 1526 | ROI2 right x value
16u | ROI2_groupx | 1528 | ROI2 amount x is binned/grouped in hw
16u | ROI2_starty | 1530 | ROI2 top y start value
16u | ROI2_endy | 1532 | ROI2 bottom y value
16u | ROI2_groupy | 1534 | ROI2 amount y is binned/grouped in hw
16u | ROI3_startx | 1536 | ROI3 left x start value
16u | ROI3_endx | 1538 | ROI3 right x value
16u | ROI3_groupx | 1540 | ROI3 amount x is binned/grouped in hw
16u | ROI3_starty | 1542 | ROI3 top y start value
16u | ROI3_endy | 1544 | ROI3 bottom y value
16u | ROI3_groupy | 1546 | ROI3 amount y is binned/grouped in hw
16u | ROI4_startx | 1548 | ROI4 left x start value
16u | ROI4_endx | 1550 | ROI4 right x value
16u | ROI4_groupx | 1552 | ROI4 amount x is binned/grouped in hw
16u | ROI4_starty | 1554 | ROI4 top y start value
16u | ROI4_endy | 1556 | ROI4 bottom y value
16u | ROI4_groupy | 1558 | ROI4 amount y is binned/grouped in hw
16u | ROI5_startx | 1560 | ROI5 left x start value
16u | ROI5_endx | 1562 | ROI5 right x value
16u | ROI5_groupx | 1564 | ROI5 amount x is binned/grouped in hw
16u | ROI5_starty | 1566 | ROI5 top y start value
16u | ROI5_endy | 1568 | ROI5 bottom y value
16u | ROI5_groupy | 1570 | ROI5 amount y is binned/grouped in hw
16u | ROI6_startx | 1572 | ROI6 left x start value
16u | ROI6_endx | 1574 | ROI6 right x value
16u | ROI6_groupx | 1576 | ROI6 amount x is binned/grouped in hw
16u | ROI6_starty | 1578 | ROI6 top y start value
16u | ROI6_endy | 1580 | ROI6 bottom y value
16u | ROI6_groupy | 1582 | ROI6 amount y is binned/grouped in hw
16u | ROI7_startx | 1584 | ROI7 left x start value
16u | ROI7_endx | 1586 | ROI7 right x value
16u | ROI7_groupx | 1588 | ROI7 amount x is binned/grouped in hw
16u | ROI7_starty | 1590 | ROI7 top y start value
16u | ROI7_endy | 1592 | ROI7 bottom y value
16u | ROI7_groupy | 1594 | ROI7 amount y is binned/grouped in hw
16u | ROI8_startx | 1596 | ROI8 left x start value
16u | ROI8_endx | 1598 | ROI8 right x value
16u | ROI8_groupx | 1600 | ROI8 amount x is binned/grouped in hw
16u | ROI8_starty | 1602 | ROI8 top y start value
16u | ROI8_endy | 1604 | ROI8 bottom y value
16u | ROI8_groupy | 1606 | ROI8 amount y is binned/grouped in hw
16u | ROI9_startx | 1608 | ROI9 left x start value
16u | ROI9_endx | 1610 | ROI9 right x value
16u | ROI9_groupx | 1612 | ROI9 amount x is binned/grouped in hw
16u | ROI9_starty | 1614 | ROI9 top y start value
16u | ROI9_endy | 1616 | ROI9 bottom y value
16u | ROI9_groupy | 1618 | ROI9 amount y is binned/grouped in hw
16u | ROI10_startx | 1620 | ROI10 left x start value
16u | ROI10_endx | 1622 | ROI10 right x value
16u | ROI10_groupx | 1624 | ROI10 amount x is binned/grouped in hw
16u | ROI10_starty | 1626 | ROI10 top y start value
16u | ROI10_endy | 1628 | ROI10 bottom y value
16u | ROI10_groupy | 1630 | ROI10 amount y is binned/grouped in hw
8s | FlatField[HDRNAMEMAX] | 1632 | Flat field file name
8s | background[HDRNAMEMAX] | 1752 | background sub. file name
8s | blemish[HDRNAMEMAX] | 1872 | blemish file name
32f | file_header_ver | 1992 | version of this file header -3 | 2.2, 3.0 | f | SPE File Version
8s | YT_Info[1000] | 1996 | Reserved for YT information
32s | WinView_id | 2996 | 0x01234567L if file created by WinX
64f | offset | 3000 | offset for absolute data scaling
64f | factor | 3008 | factor for absolute data scaling
8s | current_unit | 3016 | selected scaling unit
8s | reserved1 | 3017 | reserved
8s | string[40] | 3018 | special string for scalingx
8s | reserved2[40] | 3058 | reserved
8s | calib_valid | 3098 | flag if calibration is valid
8s | input_unit | 3099 | current input units for calib_value
8s | polynom_unit | 3100 | linear UNIT and used in the polynom_coeff
8s | polynom_order | 3101 | ORDER of calibration POLYNOM
8s | calib_count | 3102 | valid calibration data pairs
64f | pixel_position[10] | 3103 | pixel pos. of calibration data
64f | calib_value[10] | 3183 | calibration VALUE at above pos | 2.2 | d | X-Startvalue i nm
64f |  | 3199 |  | 2.2 | d | X-Stopvalue in nm
64f | polynom_coeff[6] | 3263 | polynom COEFFICIENTS
64f | laser_position | 3311 | laser wavenumber for relative WN | 2.2 | d | Laser Wavelength in nm
8s | reserved3 | 3319 | reserved
8u | new_calib_flag | 3320 | If set to 200
8s | calib_label[81] | 3321 | Calibration label (NULL termed)
8s | expansion[87] | 3402 | Calibration Expansion area
64f | offset | 3489 | offset for absolute data scaling
64f | factor | 3497 | factor for absolute data scaling
8s | current_unit | 3505 | selected scaling unit
8s | reserved1 | 3506 | reserved
8s | string[40] | 3507 | special string for scalingx
8s | reserved2[40] | 3547 | reserved
8s | calib_valid | 3587 | flag if calibration is valid
8s | input_unit | 3588 | current input units for calib_value
8s | polynom_unit | 3589 | linear UNIT and used in the polynom_coeff
8s | polynom_order | 3590 | ORDER of calibration POLYNOM
8s | calib_count | 3591 | valid calibration data pairs
64f | pixel_position[10] | 3592 | pixel pos. of calibration data
64f | calib_value[10] | 3672 | calibration VALUE at above pos
64f | polynom_coeff[6] | 3752 | polynom COEFFICIENTS
64f | laser_position | 3800 | laser wavenumber for relative WN
8s | reserved3 | 3808 | reserved
8u | new_calib_flag | 3809 | If set to 200
8s | calib_label[81] | 3810 | Calibration label (NULL termed)
8s | expansion[87] | 3891 | Calibration Expansion area
8s | Istring[40] | 3978 | special intensity scaling string
8s | Spare_6[25] | 4018
8u | SpecType | 4043 | spectrometer type (acton
8u | SpecModel | 4044 | spectrometer model (type dependent)
8u | PulseBurstUsed | 4045 | pulser burst mode on/off
32u | PulseBurstCount | 4046 | pulser triggers per burst
64f | PulseBurstPeriod | 4050 | pulser burst period (in usec)
8u | PulseBracketUsed | 4058 | pulser bracket pulsing on/off
8u | PulseBracketType | 4059 | pulser bracket pulsing type
64f | PulseTimeConstFast | 4060 | pulser slow exponential time
64f | PulseAmplitudeFast | 4068 | pulser fast exponential amplitude constant
64f | PulseTimeConstSlow | 4076 | pulser slow exponential time constant (in usec)
64f | PulseAmplitudeSlow | 4084 | pulser slow exponential amplitude constant
16s | AnalogGain | 4092 | analog gain
16s | AvGainUsed | 4094 | avalanche gain was used
16s | AvGain | 4096 | avalanche gain value
16s | lastvalue | 4098 | Always the LAST value in the header